i-STAT analyser use for patients with suspected or confirmed COVID-19

Initial preparation

- Perform hand hygiene with alcohol based hand rub and don PPE as per HPS guidance*.  This applies to staff member taking the sample and staff member transporting sample and performing analysis on i-STAT.

In the patients area

- Take sample as normal into a heparinised syringe

- Inspect the sample for air bubbles - if present, expel these by gently depressing the plunger, with a gauze swab held at the opening of the tube to catch any blood expelled.

- Cap the syringe once the air bubbles are removed

- Once sample is ready, drop it carefully into a clear plastic specimen bag held at the edge of the patient cubicle by trained i-STAT user performing analysis.  Neither the sample nor your gloves should touch the outside of the bag

- Dispose of the sharp and other parts of the packaging in appropriate clinical waste stream in the patient cubicle

Taking sample to i-STAT analyser

- The sample is now in a sealed plastic specimen bag.  This can therefore be taken to the designated i-STAT analyser.

Analysing the sample on the i-STAT

- Prior to opening the bag put the patient’s details into the i-Stat as follows:

- Press the power button to turn on i-STAT

- Press 2 for i-STAT cartridge

- Scan operator ID

- Scan patient ID

- Scan cartridge

- Position patient demographic label next to i-STAT, to denote the i-STAT is in use

Note: i-STAT should remain on a flat surface and not be moved during analysis.

- Open specimen bag

- Mix sample 8 times between the palms of your hands

- Discard 3 drops of blood to waste with a gauze swab held at the opening of the tube to catch any blood expelled

Application of Sample

This sequence must be followed:

- Place the cartridge in a cardboard drug tray

Note:If application of blood is not immediate the sample must be remixed (8 times between the palms of the hands) before analysis.

- Take the cap off the sample and place in the same drug tray.  Apply blood to CG4 or CHEM 8 cartridge and seal cartridge.

- Recap sample and place on drug tray

- Insert cartridge to begin analysis.

- Once analysis is complete discard test cartridge immediately into sharps bin. A further test cartridge may be performed at this point if required, remembering to mix the sample 8 times between the palms of the hands before application.

- Once complete discard blood syringe and cartridge into sharps bin

Note: If this sequence is not followed, there may be a risk of incorrect results being produced.

Post Procedure / Decontamination

- Remove gloves and perform hand hygiene (wash hands or hand hygiene with alcohol based hand rub)

- Using clean gloves, clean machine with Clinell sanitising wipe. The device must be cleaned prior to use on another patient

- Visible contamination - the device will need to be decontaminated immediately after use if there is visible contamination e.g. blood.  For blood apply 10, 000 ppm of chlorine solution

- Remove PPE and perform Hand Hygiene

- Once a day or as required, after decontamination, the i-Stat should be docked in MAU for upload of results to InfoHQ IT system for audit trail

* DROPLET (RESPIRATORY) PRECAUTIONS

• Disposable apron; consider fluid-resistant disposable gown if apron provides inadequate cover for the procedure/task being performed

• Disposable gloves

• Fluid-resistant Type IIR surgical face mask and goggles or full face visor if splashes anticipated